LP-SIS TUBE HEATER ONLY ---- TEST

Note:

-Tube Heater only. No Target heater

| Salts used: | Qty [µL]: |
| --- | --- |
| Li | 5 |
| Na | 5 |
| K | 5 |
| Rb | 5 |
| Cs | 5 |

Thursday, June 09, 2016, 14:50: Begin roughing.

Thursday, June 09, 2016, 15:34: Turbo pumps started.

Thursday, June 09, 2016, 15:42: Ion gauge on.

Water & RGA on and ok. Vacuum stable @ 1.9E-7T

Optics, sources, and power on.

Friday, June 10, 2016, 08:50: BIAS/EZL/EE to 30/13/1.5kV.

Friday, June 10, 2016, 08:52: Auto ramping TBHT only @ 2A min to 50A

| Heating start date and time | 2016-06-10 08:52 |
| --- | --- |
| BIAS |  |
| Voltage [kV] | 30 |
| Current [µA] | 31 |
| Extraction Electrode (EE) |  |
| Voltage [kV] | 1.5 |
| Current [µA] | 515 |
| TGHT |  |
| Current [A] | N/A |
| Voltage [V] | N/A |
| TBHT |  |
| Current [A] | 49 |
| Voltage [V] | 0.23 |

Friday, June 10, 2016, 10:02: Auto ramping TBHT only @ 1A min to 100A.

| Heating start date and time | 2016-06-10 10:02 |
| --- | --- |
| BIAS |  |
| Voltage [kV] | 30 |
| Current [µA] | 41 |
| Extraction Electrode (EE) |  |
| Voltage [kV] | 1.5 |
| Current [µA] | 513 |
| TGHT |  |
| Current [A] | N/A |
| Voltage [V] | N/A |
| TBHT |  |
| Current [A] | 99 |
| Voltage [V] | 0.55 |

Friday, June 10, 2016, 11:14: Auto ramping TBHT only @ 1A min to 150A.

Friday, June 10, 2016, 11:24: Auto Ramp halted @ 111.5A

| Heating start date and time | 2016-06-10 10:02 |
| --- | --- |
| BIAS |  |
| Voltage [kV] | 30 |
| Current [µA] | 41 |
| Extraction Electrode (EE) |  |
| Voltage [kV] | 1.5 |
| Current [µA] | 513 |
| TGHT |  |
| Current [A] | N/A |
| Voltage [V] | N/A |
| TBHT |  |
| Current [A] |  |
| Voltage [V] |  |